APPENDIX B.1

RAF PROJECT SAFETY COMMITTEE

HAZARDOUS MATERIALS AND DEVICES

(Page 1 of 4)

DO NOT WRITE IN THIS BLOCK

Project: 		___DC3____		Installation Period:	Mar 6 – Apr 30 2012

Aircraft: 		___GV____		Beginning Date: 	_8 May 2012__

Instrument Number:	_________________		Ending Date: 		_30 June 2012_

1. 	Instrument: 		CU Closed Path Laser Hygrometer (CLH-2)_

2.         Function: 		Measures water vapor and water resulting from evaporation of cloud particles

3.	Principal Investigator:	Linnea Avallone____________

Address: 		LASP, 600 UCB, Boulder, CO, 80303-0600

Telephone: 		303-492-5913__                                                                                                       4.         Instrument Operator(s): Linnea Avallone, Samuel Dorsi, Anna Luebke

5.	Is this instrument commercially produced? 	No

6.	If so, please list name and address of manufacturer:

7. 	Please list serial numbers of the instruments:

_ _N/A______

Please attach a copy of the manufacturer’s instruction manual for the device.  If this is not possible, attach a copy of those pages of the instruction manual which describe the principles of operation, hazard warnings, safety features, and safety rules.

8. 	If the instrument is not commercially produced, please provide information requested below:

Designed by: 	_Samuel Dorsi, Lars Kalnajs, Linnea Avallone __________

Organization: 	_University of Colorado Boulder______________________

Address: 		_LASP, 600 UCB, Boulder , CO 80303-0600___________

Telephone: 		_303-492-5913______________________________

Built by: 	            _Samuel Dorsi, Kraig Koski_

Organization:         _LASP, University of Colorado Boulder___________________

Address: 	            _ LASP, 600 UCB, Boulder , CO 80303-0600______________

###### APPENDIX B.1

# RAF PROJECT SAFETY COMMITTEE

HAZARDOUS MATERIALS AND DEVICES

(Page 2 of 4)

9.	Describe principles of operation, hazard warnings, safety features:

The CLH-2 uses tunable diode laser spectroscopy at 1.37 um to measure the water vapor resulting from the evaporation of cloud particles. It is fed by a heated, forward-facing inlet. Potential hazards include infrared laser light and heated inlet lines. The laser light is completely enclosed within the instrument. Heaters are thermostatically controlled and fused. Overall instrument is also fused.

10. 	If the instrument is commercially produced, has it been modified?  N/a

11. 	If modified, describe the modification.

_________________

All investigators please answer the following:

12. 	Does the instrument contain, use, or produce:

Radioactive materials		No		Compressed gases		No_____

Other ionizing radiation	No____	Non-ionizing radiation	No______

Flammable liquids		No		Laser				Yes___

Radar				No______	Flammable gases		No______

Explosive materials		No______	Toxic materials		No______

13. 	If any of the categories were checked, specify the material below (for example, amount, energy levels, physical form, etc.)

Class 3B invisible laser radiation, maximum of 10mW of CW light in the near-infrared

14. 	Please list all other chemicals you will use on board this aircraft in your experiment.

__n/a____________________________________

15.	If your experiment consumes or discharges materials, will you need to carry additional materials on board?       n/a___

###### APPENDIX B.1

# RAF PROJECT SAFETY COMMITTEE

HAZARDOUS MATERIALS AND DEVICES

(Page 3 of 4)

16. 	What and how much extra materials will you need to carry? _____N/A__________

17. 	What kind of container will you need to carry these materials? _____N/A_________

18. 	If the device utilizes a laser, please classify the laser according to ANSI Z 136.1-1973 (circle one).

Class:	III   (fully enclosed, not user serviceable)

- If your laser will be operating at a wavelength that is not eye safe,  what procedures will

be established to minimize the danger to yourself and other project participants?

The laser and its light path are completely enclosed

20.	If you are using compressed gas cylinders, what is the maximum pressure expected for each cylinder type? _n/a ___

- Will you be re-filling any compressed gas cylinders yourself, either at JeffCO or during the field deployment?___n/a__

22. 	Are there any other hazards associated with the instrument itself, the required ground support equipment or the experiment which have not so far been covered in this questionnaire?

No

23. 	How would you describe the probability of an accident resulting from the presence and use of your instrument on board the NCAR aircraft?

Low probability

###### APPENDIX B.1

RAF PROJECT SAFETY COMMITTEE

HAZARDOUS MATERIALS AND DEVICES

(Page 4 of 4)

24.	How would you describe the severity of such an accident?

Nearly all possible failures of instrument components are contained by the use of fuses to isolate electrical elements.

25. 	What precautions will you take to decrease the probability and the severity of an accident?  If any documented safety procedures from your home facility or university are available, please attach a copy of said materials to this form.

All components have been inspected and tested prior to installation on the aircraft.

_17 February 2012_____	_________________

Date				  Signature of principal investigator or operator

Linnea Avallone____

Printed name of principal investigator or operator

Reviewed by

________________________________________________

Date