Summary of Calibration Factors

GV - Basic DC3 Configuration

Cantrel et al

Pre-Calibration – 4/17/2012

Param

Variable                    Name      C1        C2       C3

Dynamic Pressure-Fuselage L. (hPa)	QCF  	   0.040	   1.000   .0

Dynamic Pressure-Radome (hPa)	QCR 	  -0.032     0.999	 .0

Total Temp: Rosemont Un-Heated (C)	TTRL	 -81.781    21.082   .251

Total Temp: HARCO Heated #1 (C)	TTHR1 	 -79.383    21.175   .271

Total Temp: HARCO Heated #2 (C)      TTHR2   -78.554    21.118   .256

Vertical Press. Diff. Radome(mb)	ADIFR	   0.001     0.998   .0

Horizontal Press. Diff. Gustpod(mb) 	BDIFR     0.060     1.0014  .0

Vertical Press. Diff. Gustpod(mb)	ADIFR_GP  0.0      -1.0     .0

Horizontal Press. Diff. Radome(mb) 	BDIFR_GP  0.0       1.0     .0

Aircraft Pitch Angle (deg)	PITCH+	   0.50      1.0     .0

Static Pressure-Fuselage (hPa)	PSF	  -0.148     1.0001  .0

Dew Point Sensor Left Serial (C)	DPL	   0.0       1.0     .0

Dew Point Sensor Right Serial (C)	DPR	   0.0       1.0     .0

Cabin Pressure (hPa)	PCAB     -3.182   108.82   -.0181

Cavity Pressure, Lft Dew Point (mb)	CAVP_DPL  1.577   219.15    .0625

Cavity Pressure, Rgt Dew Point (mb)	CAVP_DPR  1.081   219.43    .0048

# +  Corrected for systematic offset

# Calibration Summary

page 2

Post Calibration – 7/10/2012

Param

Variable                    Name      C1        C2       C3__

Dynamic Pressure-Fuselage L. (hPa)	QCF  	   0.190	   1.000   .0

Dynamic Pressure-Radome (hPa)	QCR 	  -0.182	   1.001	 .0

Total Temp: Rosemont Un-Heated (C)	TTRL*	 -84.004    22.314   .3027

Total Temp: HARCO Heated #1 (oC)	TTHR1*	 -82.117    22.593   .3054

Total Temp: HARCO Heated #2 (C)      TTHR2*  -82.217    22.802   .2819

Vertical Press. Diff. Radome(mb)	ADIFR	   0.006     0.9987  .0

Horizontal Press. Diff. Gustpod(mb) 	BDIFR     0.027     0.9973  .0

Vertical Press. Diff. Gustpod(mb)	ADIFR_GP  0.0      -1.0     .0

Horizontal Press. Diff. Radome(mb) 	BDIFR_GP  0.0       1.0     .0

Aircraft Pitch Angle (deg)	PITCH	   0.30      1.0     .0

Static Pressure-Fuselage (hPa)	PSF	  -0.341     1.0018  .0

Dew Point Sensor Left Serial (C)	DPLS	     not functional

Dew Point Sensor Right Serial (C)	DPRS	  -0.507     0.986   .0002 Cabin Pressure (hPa)	PCAB     -2.873   108.93   -.0016

Cavity Pressure, Lft Dew Point (mb)	CAVP_DPL  1.660   219.68    .0249

Cavity Pressure, Rgt Dew Point (mb)	CAVP_DPR  2.070   219.69    .0048

Cn Counter Cavity Pressure (mb)	PCN      -2.197   109.02   -.342

* Post cal used in final processing.

Calibration Maneuvers –

RAF uses sets of in-flight maneuvers specifically designed to help evaluate the performance of the 3-D wind measurement system on the aircraft.  Sets of maneuvers were flown on one test flight and one research flight: tf01, rf05, rf09, rf10, rf22 and FF03 respectively. Systematic offsets to THDG (0.4 deg) and PITCH (0.5 deg) were determined and have been applied for final data processing.

PMS - 1-D Probe Bin Sizes

CDP probe 058 was used on the project.  The minimum droplet size (microns) for each sampling bin are listed below:

CDP

bin#1 		 3.0 um

bin#2		 	 4.0

bin#3			 5.0

bin#4			 6.0

bin#5			 7.0

bin#6			 8.0

bin#7			 9.0

bin#8			10.0

bin#9			11.0

bin#10		12.0

bin#11		13.0

bin#12		14.0

bin#13		16.0

bin#14		18.0

bin#15		20.0

bin#16		22.0

bin#17		24.0

bin#18		26.0

bin#19		28.0

bin#20		30.0

bin#21		32.0

bin#22		34.0

bin#23		36.0

bin#24		38.0

bin#25		40.0

bin#26		42.0

bin#27		44.0

bin#28		46.0

bin#29		48.0

bin#30		50.0

bin#31		52.0