12/09/2011

Data Quality Report

Operation IceBridge 2011 - GV

Studinger et al

This summary has been written to outline basic instrumentation problems affecting the quality of the data set and is not intended to point out every bit of questionable data. It is hoped that this information will facilitate use of the data as the research concentrates on specific flights and times.  The following report covers only the RAF supplied instrumentation.

Section I: General Discussion

RAF staff have reviewed the data set for instrumentation problems.  When an instrument has been found to be malfunctioning, specific time intervals are noted.  In those instances the bad data intervals have been filled in the netCDF data files with the missing data code of -32767. In some cases a system will be out for an entire flight.  Virtually all measurements made on the aircraft require some sort of airspeed correction or the systems simply do not become active while the aircraft remains on the ground.  None of the data collected while the aircraft is on the ground should be considered as valid.

SPECIAL NOTE: RAF flies redundant sensors to assure data quality. Performance characteristics differ from sensor to sensor with certain units being more susceptible to various thermal and dynamic effects than others.  Good comparisons were typically obtained between the two static pressures (PSFC, PS_A), the four temperature sensor outputs (ATFH1, ATFH2, ATHR1, ATHR2), the three dynamic pressures (QCRC, QCFC, QC_A), and the two dew pointers (DPLS,DPRS). The primary research pressure system uses inputs from the radome gust pod.  On the flight (rf08) where the radome measurements were affected by cloud penetrations, PS_A was substituted as the reference pressure (PSXC) on all with only a minor decrease in overall accuracy. Similarly, QC_A was used for QCXC and TAS_A was used for TASX.

Position Data. A Novatel Global Positioning System (GGPS) was used as an accurate position reference during the program.  The GPS data were updated with real-time Omnistar XP signals.  The system generally performed well.  It is recommended that the GPS data be used as the position reference (GGLAT, GGLON). There may be occasional spikes or discontinuous shifts in these values due to satellite geometry and aircraft maneuvering. The algorithm referred to in the 3D-Wind Data section below also blends the GPS and IRS position data to yield a best position (LATC, LONC) that generally removes the GPS spikes.

3D-Wind Data. The primary wind data for this project were derived from measurements taken with the radome wind gust package.  As is normally the case with all wind gust systems, the ambient wind calculations can be adversely affected by either sharp changes in the aircraft's flight attitude or excessive drift in the onboard inertial reference system (IRS).  Turns, or more importantly, climbing turns are particularly disruptive to this type of measurement technique.  Wind data reported for these conditions should be used with caution.

Special sets of in-flight calibration maneuvers were conducted to aid in the performance analysis of the wind gust measurements.  The calibration data identified a systematic bias in the pitch and sideslip parameters. These offsets have been removed from the final data set. Drift in the IRS accelerometers are removed using an algorithm that employs a complementary high-pass/low-pass filter that removes the long term drift with the accurate GPS reference and preserves the shorter term fluctuations measured by the IRS. Both the GPS corrected and basic uncorrected values are included in the final data set for the purpose of data quality review.  RAF strongly recommends that the GPS corrected inertial winds be used for all research efforts (WSC,WDC,UXC,VYC,WIC,UIC,VIC).

Ambient Temperature Data. Temperature measurements were made using the standard heated HARCO sensor (ATHR1, ATHR2) and a heated Rosemount temperature sensor (ATFH1, ATFH2).  In addition there is an output from a second Rosemount sensor that feeds the avionics (AT_A).  While the four research outputs generally agreed under most conditions, there is a significant difference between those systems and the avionics sensor at very low temperatures.  Comparisons against alternate references (GPS dropsondes) and an evaluation using derived dependent variables (UXC, VYC, WSC, WDC) have led to the conclusion that the avionics temperature is more accurate.  For this reason TT_A and AT_A have been used as the values for TTX and ATX, respectively.

Humidity Data. Humidity measurements were made using two thermoelectric dew point sensors. A comparison of the dew point data (DPLC, DPRC) yielded good correlation in instrument signatures during the largest portions of the flights when both instruments were functioning normally.  Under conditions where the units had been cold soaked at high altitude or experienced a rapid transition into a moist environment, both units showed a tendency to overshoot or lag in time by multiple seconds. This problem can also result in loss of data.  Even at their best, the response of the thermoelectric dew point sensors is roughly 2 seconds.  Response times are dependent upon ambient dew point depression and can exceed 10-15 seconds under very dry conditions.   During the rapid climbs and descents characterizing the beginning and ending of the project flights, the slow time response resulted in unrealistic super-saturation dew points.  These data have been left in the data sets to provide an indication of vertical dew point trends as opposed to accurate values at specific altitudes. The extensive high altitude work during the project placed the units under stress with extreme dew point depressions. Under these conditions overall accuracy is reduced. Based on overall performance DPLC was selected as the reference signal (DPXC).

Altitude Data. The altitude of the aircraft was measured in several ways. A pressure based altitude (PALT,PALTF) is derived from the static pressure using the hydrostatic equation and the U.S. Standard Atmosphere, which assumes a constant surface pressure of 1013mb and a surface temperature of 288.15 K.

The GPS positioning system also provides an altitude readout (GGALT).  This output normally provides a fairly accurate MSL altitude based on an ellipsoid model of the Earth (WGS-84). The RAF recommends the use of the GPS altitude as the reference value for this measurement.

*	*	*	*	*	*	*	*	*	*	*	*